## Cover Page
Intented Use of 400-013435-00 distortion_pixel_align_calc spreadsheet | The scope_dov spreadsheet is used for design verification testing of the Ottava System (specifically the Laparoscope Direction of View (DOV) requirements). The excel spreadsheet is used to calculate the direction of view of the Laparoscope based on the checkerboard target analysis from Imatest.
DOV Tab | Calculates the DOV based on two inputs - (1) pixel coordinates of the checkerboard target image center captured from the left and right eye of the Laparoscope and (2) the magnification in pixel of both left and right eye of the Laparoscope.
Instructions For Use | 1. Enter the pixel coordinates of the image center indicated in the "Input" section of the table in the DOV tab.
2. Enter the magnification (in pixel) indicated in the "Input" section of the table in the DOV tab.
3. Obtain the DOV error value at the bottom of the sheet.

## DOV
Input - Target Center at Sensor
Working Distance (mm) | Left Eye |  |  | Right Eye
| Image Target Center Coordinate |  | Magnification (Pixel) | Image Target Center Coordinate |  | Magnification (Pixel)
| X | Y |  | X | Y
35
50
Calculations
1. Convert Target Center from Pixel to Distance
Raw Pixel Size (mm) |  |  | 0.00224
Raw Vertical Pixel Number |  |  | 720
Remapped Vertical Pixel Number |  |  | 1080
Remapped Pixel Size (mm) |  |  | 0.0014933
Remapped Resolution (H x V) |  |  | 1920 |  | 1080
Checkerboard Target Square Size (mm) |  |  | 4
Working Distance (mm) | Left Eye |  |  | Right Eye
| Image Target Center Coordinate |  | Magnification (Distance) | Image Target Center Coordinate |  | Magnification (Distance)
| X | Y |  | X | Y
35 | 960 | 540 | 0 | 960 | 540 | 0
50 | 960 | 540 | 0 | 960 | 540 | 0
2. Convert from Sensor Distance to Object Distance
Working Distance (mm) | Left Eye |  | Intentionally Left Blank | Right Eye |  | Intentionally Left Blank
| Image Target Center Coordinate |  |  | Image Target Center Coordinate
| X | Y |  | X | Y
35 | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
50 | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
Slope | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
3. Direction of View Error
| X | Y | Overall
DOV Error (Deg) | #DIV/0! | #DIV/0! | #DIV/0!
Performed by: __________________________ |  |  |  | Date: __________________________
Reviewed by: __________________________ |  |  |  | Date: __________________________